## FormTitan
Header
Fail Analisis | 230125(2)
Internal Quality Control (IQC) ID: | IQC POW 220125
Reagent Blank (RB) ID: | RB POW 220125
Arsenic (As)
RB (ppb): | 0.028
Sample weight IQC A | 0.505
Sample weight IQC B | 0.506
Blank Conc. (ppb) | 54.49
IQC POW A (ppb): | 4742.819
IQC POW B (ppb): | 4735.762
Cadmium(Cd)
RB (ppb): | 0.021
Sample weight IQC A | 0.505
Sample weight IQC B | 0.506
Blank Conc. (ppb) | 18.023
IQC POW A (ppb): | 296.581
IQC POW B (ppb): | 306.292
Mercury (Hg)
RB (ppb): | 0.66
Sample weight IQC A | 0.505
Sample weight IQC B | 0.506
Blank Conc. (ppb) | 41.137
IQC POW A (ppb): | 538.04
IQC POW B (ppb): | 548.145
Plumbum(Pb)
RB (ppb): | 1.097
Sample weight IQC A | 0.505
Sample weight IQC B | 0.506
Blank Conc. (ppb) | 238.897
IQC POW A (ppb): | 10196.046
IQC POW B (ppb): | 10437.1
Nama dan Tarikh
Nama | AMIR / MAISARAH
Tarikh | 2025-01-23

## FormGH
Header
Fail Analisis | 210125(1)
Internal Quality Control (IQC) ID: | IQC LIQ 130125
Reagent Blank (RB) ID: | RB GH A 130125
Arsenic (As)
RB (ppb): | 0.024
Sample weight IQC A | 1.506
Sample weight IQC B | 1.505
Blank Conc. (ppb) | 13.335
IQC POW A (ppb): | 4623.517
IQC POW B (ppb): | 4751.689
Cadmium(Cd)
RB (ppb): | 0.019
Sample weight IQC A | 1.506
Sample weight IQC B | 1.505
Blank Conc. (ppb) | 4.079
IQC POW A (ppb): | 288.316
IQC POW B (ppb): | 287.882
Mercury (Hg)
RB (ppb): | 0.503
Sample weight IQC A | 1.506
Sample weight IQC B | 1.505
Blank Conc. (ppb) | 7.173
IQC POW A (ppb): | 488.44
IQC POW B (ppb): | 493.932
Plumbum(Pb)
RB (ppb): | 1.553
Sample weight IQC A | 1.506
Sample weight IQC B | 1.505
Blank Conc. (ppb) | 476.841
IQC POW A (ppb): | 9803.191
IQC POW B (ppb): | 9721.95
Nama dan Tarikh
Nama | AMIR / MAISARAH
Tarikh | 2025-01-21

## IQC TITAN
Unit Analisis Tradisional
Laporan Internal Quality Control ICPMS
Fail Analisis: |  |  | 230125(2)
Internal Quality Control (IQC) ID: |  |  | IQC POW 220125
Reagent Blank (RB) ID: |  |  | RB POW 220125
ARSENIC |  |  |  |  | RB (ppb): | 0.028 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of As in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<6%)
IQC Sample A | 0.505 | 54.49 | 4742.819 | 4688.329 | 94.7% | Within Warning Limit □ |  | 0.1506%
Within Control Limit □
IQC Sample B | 0.506 | 54.49 | 4735.762 | 4681.272 | 94.75% | Within Warning Limit □
Within Control Limit □
CADMIUM |  |  |  |  | RB (ppb): | 0.021 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of Cd in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<9%)
IQC Sample A | 0.505 | 18.023 | 296.581 | 278.558 | 93.78% | Within Warning Limit □ |  | 3.426%
Within Control Limit □
IQC Sample B | 0.506 | 18.023 | 306.292 | 288.269 | 97.24% | Within Warning Limit □
Within Control Limit □
MERCURY |  |  |  |  | RB (ppb): | 0.66 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of Hg in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<8%)
IQC Sample A | 0.505 | 41.137 | 538.04 | 496.903 | 100.4% | Within Warning Limit □ |  | 2.013%
Within Control Limit □
IQC Sample B | 0.506 | 41.137 | 548.145 | 507.008 | 102.6% | Within Warning Limit □
Within Control Limit □
PLUMBUM |  |  |  |  | RB (ppb): | 1.097 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of Pb in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<5%)
IQC Sample A | 0.505 | 238.897 | 10196.046 | 9957.149 | 100.6% | Within Warning Limit □ |  | 2.392%
Within Control Limit □
IQC Sample B | 0.506 | 238.897 | 10437.1 | 10198.203 | 103.2% | Within Warning Limit □
Within Control Limit □
Tandatangan Penganalisa & tarikh |  |  |  | AMIR / MAISARAH |  |  | 2025-01-23
Tandatangan Penyemak & tarikh

## IQC GH
Unit Analisis Tradisional
Laporan Internal Quality Control ICPMS
Fail Analisis: |  |  | 210125(1)
Internal Quality Control (IQC) ID: |  |  | IQC LIQ 130125
Reagent Blank (RB) ID: |  |  | RB GH A 130125
ARSENIC |  |  |  |  | RB (ppb): | 0.024 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of As in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<6%)
IQC Sample A | 1.506 | 13.335 | 4623.517 | 4610.182 | 92.57% | Within Warning Limit □ |  | 2.742%
Within Control Limit □
IQC Sample B | 1.505 | 13.335 | 4751.689 | 4738.354 | 95.08% | Within Warning Limit □
Within Control Limit □
CADMIUM |  |  |  |  | RB (ppb): | 0.019 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of Cd in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<9%)
IQC Sample A | 1.506 | 4.079 | 288.316 | 284.237 | 95.12% | Within Warning Limit □ |  | 0.1528%
Within Control Limit □
IQC Sample B | 1.505 | 4.079 | 287.882 | 283.803 | 94.92% | Within Warning Limit □
Within Control Limit □
MERCURY |  |  |  |  | RB (ppb): | 0.503 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of Hg in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<8%)
IQC Sample A | 1.506 | 7.173 | 488.44 | 481.267 | 96.64% | Within Warning Limit □ |  | 1.135%
Within Control Limit □
IQC Sample B | 1.505 | 7.173 | 493.932 | 486.759 | 97.68% | Within Warning Limit □
Within Control Limit □
PLUMBUM |  |  |  |  | RB (ppb): | 1.553 |  | [MDL:50 ppb]
| Sample weight (g) | Blank Conc. (ppb) | Found Conc. (ppb) | Conc of Pb in QC samples | % Recovery (80 - 115%) |  |  | %RPD (NMT<5%)
IQC Sample A | 1.506 | 476.841 | 9803.191 | 9326.35 | 93.64% | Within Warning Limit □ |  | 0.8749%
Within Control Limit □
IQC Sample B | 1.505 | 476.841 | 9721.95 | 9245.109 | 92.76% | Within Warning Limit □
Within Control Limit □
Tandatangan Penganalisa & tarikh |  |  |  | AMIR / MAISARAH |  |  | 2025-01-21
Tandatangan Penyemak & tarikh